- Purpose

This document specifies all interfaces for the Fiber Frequency Comb modular system, including those for the Comb Module, the Pump Laser Module, and the FPGA Controller.

References

50-00038-R01 Pump Laser Module.pdf

50-00054-R01 FPGA Controller.pdf

50-00068-R01 Comb Module.pdf

FFC-003 Performance Specifications

Definitions & Abbreviations

| Term | Definition |
| --- | --- |
| API | Application Programming Interface |
| CW | Continuous Wave |
| fCEO | Carrier-Envelope Offset Frequency. Detected as the RF signal output from an f-2f interferometer. Typically, this signal is phase-locked to a stable RF reference signal. |
| fOPT | Optical Frequency. Detected as the RF heterodyne beat signal between an input CW reference laser and one of the FFC comb modes. Typically, this signal is phase-locked to a stable RF reference signal. |
| fREP | Repetition Rate. Functions as the RF signal output from an optical clock. |
| FFC | Fiber Frequency Comb |
| MCU | Microprocessor Control Unit |
| PM | Polarization Maintaining |
| PZT | Piezoelectric Transducer |
| RF | Radiofrequency |
| SC | Supercontinuum |

System Description

The system provides a fully stabilized FFC capable of performing optical frequency division of an optical clock. Two versions of the system will be provided: (a) one that can divide an input 1556 nm optical clock signal and (b) the other capable of dividing an input 778 nm optical clock signal. The comb system comprises three modules:

Comb Module

Detects fCEO of the FFC via the “f-2f” self-referencing technique.

Detects fREP output of the comb oscillator.

Detects optical heterodyne beat signal fOPT between a (filtered) comb mode and an input CW optical frequency at either 1556 nm or 778 nm.

Provides a PZT transducer for fast tuning of fREP (fOPT), including a built-in high-voltage driver.

Provides precise and stable temperature control of the laser cavity for coarse tuning of fREP (fOPT).

Outputs a portion of the optical output of the mode-locked oscillator.

Outputs a portion of the optical output of the supercontinuum.

Controlled either by I2C, USB, or GPIO (TTL-level logic). The GPIO can be configured for UART communications.

Pump Laser Module

Provides for up to four pump lasers to pump the erbium gain fiber of the FFC. The 1556 nm System requires two pump lasers. The 778 nm System requires three pump lasers.

Allows full control of each pump laser’s injection current and temperature via I2C.

Provides a current modulation input for electronically tuning fCEO.

FPGA Controller

Phase locks an input fCEO signal to a user-specified RF reference frequency. Accepts an input fCEO signal and digitally processes it to generate an output control signal to the Current Modulation port of the Oscillator Pump Laser.

Phase locks an input fOPT signal to a user-specified RF reference frequency. Accepts an input fOPT signal and digitally processes it to generate an output control signal to the PZT Modulation port of the Comb Module.

Controlled via an application program running on an external PC.

Capable of referencing the phase-lock offset frequencies to an external clock.

System Block Diagram

Figure 1: Block diagram of the comb system showing the three Vescent components in blue (Comb Module, Pump Laser Module, and FPGA Controller) and their interconnections with external user system components. Power connections are red; computer control connections are gold; optical fiber connections are green; and analog electrical signals are grey. Numbered labels correspond to IDs referred to in the connector specifications given below for each module.

Comb Module

Connector Specifications (see Figure 2 for panel locations)

| ID | Function | Signals | Module Connector Specification |
| --- | --- | --- | --- |
| 1 | PZT Mod Input | Analog Voltage (0-8 VDC) GND | SMA, Female |
| 2 | Digital Control | I2C Serial Interface (SDA, SCL) GPIOx3 (UART capable) | TE Connectivity 5747844-5 Connector, D-Subminiature, 9 Socket, Rear Panel Mount 90˚ Receptacle, Through Hole Mating Cable Assembly: Conec 15-008163 (Mouser P/N 706-15-008163) Cable, DSUB, 9 Pin Plug, Unshielded, 22AWG, 4M (Might also consider a shielded cable) |
| 3 | DC Power Input | +5 VDC (< 1.5 A) +5 VDC Return +15 VDC (< 0.25 A) -15 VDC (< 0.25 A) ±15 VDC Return | Conec 43-01211 Connector, M12, 5 Pin, Front Panel Mount 90˚ Receptacle, Through Hole Mating Cable Assembly: TE Connectivity T4161320005-002 (Digi-Key P/N A142281-ND) Cable, M12, 5 Socket Plug, Shielded, 22AWG, 1M |
| 4 | fCEO Output | RF (100 MHz BW, -30 dBm) GND | SMA, Female |
| 5 | fOPT Output | RF (100 MHz BW, -10 dBm) GND | SMA, Female |
| 6 | fREP Output | RF (500 MHz, +5 dBm) GND | SMA, Female |
| 7 | Oscillator Pump Input | PM980 Optical Fiber Input | LC/APC |
| 8 | Amplifier 1 Pump Input | PM980 Optical Fiber Input | LC/APC |
| 9 | Amplifier 2 Pump Input | PM980 Optical Fiber Input (used for 788 nm version only) | LC/APC |
| 10 | CW Laser Input | PM1550/PM780 Optical Fiber Input | LC/APC |
| 11 | Oscillator Light Output | PM1550 Optical Fiber Output | LC/APC |
| 12 | SC Light Output | PM1550 Optical Fiber Output | LC/APC |
| 13 | USB Serial Port | USB 2.0 | Molex 0548190572 Connector, USB Mini B 2.0, 90˚ Receptacle, SMT |

Figure 2: Panel connections for the Comb Module. Top: Front Panel; Bottom: Rear Panel.

Digital Control Connector Pinout (TE Connectivity 5747844-5)

Dimensions in inches.

| Pin | Signal |  | Pin | Signal |
| --- | --- | --- | --- | --- |
| 1 | RESERVED |  | 6 | RESERVED |
| 2 | GPIO_A |  | 7 | GPIO_C |
| 3 | GPIO_B |  | 8 | GND |
| 4 | GND |  | 9 | I2C_SCL |
| 5 | I2C_SDA |  |  |  |

DC Power Input Connector Pinout (Conec 43-01211)

Dimensions in millimeters.

| Pin | Signal |
| --- | --- |
| 1 | +5V_RTN* |
| 2 | +5V |
| 3 | +15V |
| 4 | -15V |
| 5 | 15V_RTN* |

* Returns are galvanically connected.

Mechanicals

See drawing 50-00068-R01 Comb Module for enclosure dimensions and mounting hole specifications.

GPIO (GPIO_A, GPIO_B, GPIO_C)

All GPIO pins operate with +3.3 V CMOS logic.

GPIO may be configured for different functionality, including:

System diagnostics (e.g., FAULT, PWR_GOOD, MODE_LOCK_GOOD)

System commands (e.g., RESET, STANDBY)

UART communications (ask Vescent for additional configuration information)

Exact GPIO functions are TBD.

Power Sequencing

All power rails (+5 V /+15 V / -15 V) shall be enabled simultaneously.

I2C Communications (I2C_SDA, I2C_SCL)

All I2C pins operate with +3.3 V logic.

I2C address is TBD.

The Comb Module is configured to be a Slave I2C device.

The I2C Clock (I2C_SCL) shall operate at a frequency of 400 kHz.

The Comb Module shall provide pull-up resistors to +3.3 V for the I2C lines.

Software Interface

Software may be used to control and configure the temperature controller and the high-voltage amplifier, as well as read certain states of the module. Alternatively, the Comb Module can be configured by software for control by digital logic via the three GPIO pins.

I2C: Complete I2C command list will be provided by Vescent.

USB and UART: Complete text-based API will be provided by Vescent.

Pump Laser Module

Connector Specifications (see Figure 3 for board locations)

| ID | Function | Signals | Board Connector Specification |
| --- | --- | --- | --- |
| 1 | Pump 1 Output (Oscillator) | 980 nm optical pump, <300 mW | PM980 pigtail, >2 m, LC/APC terminated |
| 2 | Pump 2 Output (Amplifier 1) | 980 nm optical pump, <1000 mW | PM980 pigtail, >2 m, LC/APC terminated |
| 3 | Pump 3 Output (Amplifier 2, 778 nm only) | 980 nm optical pump, <1000 mW | PM980 pigtail, >2 m, LC/APC terminated |
| 4 | Oscillator Current Modulation Input (fCEO control) | Analog Voltage (-10 to +10 VDC) GND | SMA, Female |
| 5 | Fan Power Outputs (x2) (Optional Use) | +5 VDC +5 VDC Return | TE Connectivity MTA100 Series Header, 2 pin, 0.1” Reference Designator(s): J29, J30 |
| 6 | I2C Serial Interface | I2C Serial Interface (SDA, SCL) GND | TE Connectivity MTA100 Series Header, 4 pin, 0.1” Reference Designator(s): J26 |
| 7 | Interlock (Digital Controls) | Interlock Power Good (Not used) Fault (Not used) System Reset (Not used) | TE Connectivity MTA100 Series Header, 6 pin, 0.1” Reference Designator(s): J27 |
| 8 | DC Power Input | +5 VDC (< 1.5 A) +5 VDC Return +15 VDC (< 0.15 A) -15 VDC (< 0.15 A) ±15 VDC Return | TE Connectivity MTA100 Series Header, 7 pin, 0.1” Reference Designator(s): J28 |

Figure 3: Board connections to the Pump Laser Module.

Fan Power Connector Pinouts (J29, J30)

| Pin | Signal |
| --- | --- |
| 1 | V+ (+5 VDC) |
| 2 | V- (+5 VDC Return) |

I2C Connector Pinouts (J26)

| Pin | Signal |
| --- | --- |
| 1 | I2C_SDA |
| 2 | GND |
| 3 | I2C_SCL |
| 4 | GND |

Digital Interface Connector Pinouts (J27)

| Pin | Signal |
| --- | --- |
| 1 | PWR_GOOD |
| 2 | FAULT |
| 3 | SYS_RESET |
| 4 | GND |
| 5 | INTERLOCK_EN |
| 6 | GND |

Power Connector Pinouts (J28)

| Pin | Signal |
| --- | --- |
| 1 | +5V |
| 2 | +5V |
| 3 | +5V_RTN* |
| 4 | +5V_RTN* |
| 5 | +15V |
| 6 | 15V_RTN* |
| 7 | -15V |

* These returns are galvanically connected.

  - Mechanicals

    - See drawing 50-00038-R01 Pump Laser Module for board dimensions and mounting hole specifications.

  - Forced Air Cooling

    - The Pump Laser Module requires about 17 cfm of forced air directed across the module’s cooling fins to convectively cool the electronics and lasers.

    - Two on-board +5 VDC power connectors (J29 and J30) are available to power external fans. +5 V power is enabled when ±15 V power rails are applied to the Pump Laser Module.

  - Power Sequencing

    - All power rails (+5 V /+15 V / -15 V) shall be enabled simultaneously.

Laser Interlock (INTERLOCK_EN)

A safety interlock is provided by the INTERLOCK_EN pin of the Digital Controls Connector (J27).

No current will flow to any of the pump laser diodes when INTERLOCK_EN is low.

The INTERLOCK signal operates with +3.3 V CMOS levels.

Digital Status Signals (PWR_GOOD, FAULT, SYS_RESET)

Not usable in this implementation.

These pins can be left floating.

I2C Communications (SDA, SCL)

All I2C pins operate with +3.3 V logic.

I2C address is TBD.

The Pump Laser Module is configured to be a Slave I2C device.

The I2C Clock (I2C_SCL) shall operate at a frequency of 400 kHz.

The Comb Module shall provide pull-up resistors to 3.3 V for the I2C lines.

Software Interface

The temperature and injection current of each pump laser can be fully controlled via an I2C command set to be provided by Vescent.

FPGA Controller

Connector Specifications (see Figure 4 for panel locations)

| ID | Function | Signals | Panel Connector Specification |
| --- | --- | --- | --- |
| 1 | fCEO Input | RF (50 MHz BW, -30 to +10 dBm) GND | SMA, Female |
| 2 | Current Modulation Output (fCEO control) | Analog Voltage (0-8 VDC) GND | SMA, Female |
| 3 | fOPT Input | RF (50 MHz BW, -30 to +10 dBm) GND | SMA, Female |
| 4 | PZT Modulation Output (fOPT control) | Analog Voltage (0-8 VDC) GND | SMA, Female |
| 5 | Reference Clock Input | RF (200 MHz, +10 dBm) GND | SMA, Female |
| 6 | USB Serial Port | USB 2.0 | USB Standard Type B |
| 7 | AC Wallplug Power Input | 120 VAC / 60 Hz | IEC C14 Inlet |

Figure 4: Panel connections to the FPGA Controller. Left: Front Panel; Right: Rear Panel.

Mechanicals

See drawing 50-00054-R01 FPGA Controller for enclosure dimensions.

Reference Clock Input

The FPGA Controller can be synchronized to an external reference clock that is input to Connection #5 SMA located on the Rear Panel.

The reference clock input must be 200 MHz at +10 dBm.

It should be possible to condition the second harmonic of the Comb Module repetition rate output at 100 MHz as an input reference clock signal. Consult Vescent if there is interest in doing this.

Software Interface

Operation of the FPGA Controller requires a software application running on a PC Windows computer via the USB 2.0 interface. The software is open-source and will be provided by Vescent.

The FPGA open-source software provides a graphical user interface for:

Controlling the PIID loop filter parameters for the phase locks. (Vescent will recommend operating values in a provided Certificate of Conformance.)

Enabling the phase locks.

Setting digital input filtering and lock frequencies.

Examining phase-lock diagnostics, including phase noise measurements.

Change History

| Release Date | Revision | Description of Changes |
| --- | --- | --- |
| 08-06-2021 | 01 | Initial Release |